## Slide 1: Laser-Induced Breakdown Spectroscopy (LIBS) for the RAPID Elemental analysis of solids
C&E News Webinar

## Slide 2: Outline of today’s seminar
Overview – What is LIBS?
  LIBS / Hardware overview
  Principle of operation
  Strengths of LIBS as an analytical technique
Example applications
  Microanalysis – Glass, Gems, Forensics
  Bulk analysis – Aluminum, Steel
  Complex samples – Geological
  Depth profiling / surface mapping – Thin Films
TSI’s ChemReveal LIBS Desktop analyzer
Summary

## Slide 3: Typical LIBS Hardware schematic
Line position provides species identification.
Line intensity provides species concentration.
Laser energy used for both ablation and plasma formation
Pulsed
laser,
video
Fiber optic
Atomic
emission
collection
Laser
plasma
Sample
Spectrometer
and detector

## Slide 4: Elements across the periodic table can be detected with LIBS
Approximate DL
1-10 ppm
10-100 ppm
10-200 ppm
0.05-0.5%
Notes: Most elements can be measured with LIBS

## Slide 5: Substantial and growing LIBS interest

## Slide 6: LIBS has several “steps”
ABLATION – the laser removes / vaporizes material from the sample
PLASMA – electrons are stripped from the atoms by the high electric field, followed by absorption and multiphoton ionization to create a plasma
EXCITATION – the plasma electronically excites the vaporized material – forming excited atoms
EMISSION – as the plasma cools, atoms return to the ground state and give off LIGHT
Ground
Excited states

## Slide 7: Overview of LIBS plasma emission
Plasma emission is a function of time, as plasma expands and cools
Emission in the early time is dominated by white light – “Brehmstraalung emission” from electrons
Collect emission on a delay and we can see emission from ions, atoms, and molecules
0
1 s
10 s
Local thermodynamic
Equilibrium
Ionic emission
0.25-3 s
Atomic emission
0.5-50 s
Molecular recombination
2-100 s
Time
Ablation

## Slide 8: “Matrix” dependence
Differing results with different types of material – common to almost all methods
Primarily due to ablation
  Ablation on various materials is of differing effectiveness
  Controls amount of analyte vaporized
  Directly affects signal and energy in plasma
Once aerosols are formed, process is similar for most materials
craters
Aerosols,
Fractionation,
Matrix
Implication  LIBS measurements (detection limits, calibrations, etc.) are specific to a particular class of materials

## Slide 9: Implementation of LIBS
Qualitative identification can be done on any material
Different system components yield different performance (material detection limits)
Each type of material needs its own “calibration curve” for quantification, because of matrix effects
Typical detection limits ~10 ppm
RSD ~ 2-3%

## Slide 10: Primary benefits of LIBS
Rapid analysis / No reagents
“Stand-off” projection of laser
Light elements (Z<12) that XRF cannot see
Micro- to Macro-Analysis
Surface mapping and depth profiling
Compliment ICP-OES, AA, ICP-MS with rapid screening
No need to modify samples
Connect the dots to directly measure C, N, O, Be, Li, B, etc.
Analysis spot sizes from 10 μm to 400 μm
Evaluate coatings and surfaces without the complexity of GD or SEM

## Slide 11: Poll question #1
Which characteristic of LIBS would be most valuable in your applications?
Direct analysis – no sample prep
Light elements
Micro-analysis
Surface mapping
Depth Profiling
Other

## Slide 12: Example applications
What can LIBS do?

## Slide 13: Example 1: Investigation – Sample analysis occurs with little damage
Can you see the analysis points on this gemstone?

## Slide 14: Small analysis spots yield many advantages
Expansion of the view shows that this is a 25-micron spot
Not distinguishable under jeweler’s inspection loupe
Good for analysis of small features
Range of spot sizes can span from micro- to macro-analysis
Low resolution
High resolution

## Slide 15: Approximate DL in USGS Basaltic Glasses with LIBS
Element | LOD (µg/g) | Element | LOD (µg/g) | Element | LOD (µg/g)
Alkali and Alkali Earth |  | Transition Metals |  | Lanthanides/Rare Earth Elements | 
Li | 5.66 | Sc | 41.2 | La | 164
Be | 2.98 | Ti | 136 | Ce | 261
Rb | ND | V | 36.0 | Pr | 476
Sr | 25.6 | Cr | 24.9 | Nd | 92.7
Cs | ND | Mn | 191 | Sm | 512
Ba | 34.9 | Fe | Major Element | Eu | 337
 |  | Co | 46.7 | Gd | 390
 |  | Ni | 17.7 | Tb | 525
 |  | Cu | 6.59 | Dy | 359
 |  | Zn | 59.6 | Ho | 577
 |  |  |  | Er | ND
 |  |  |  | Tm | ND
 |  |  |  | Yb | 4.85
 |  |  |  | Lu | 269

## Slide 16: Features of LIBS that Glass, Forensic, and Gem analyses require
Minimally-destructive microanalysis
No sampling / digestion
Targeting to measure composition of particular features
High sensitivity to particular light elements such as Be and Li
 The ChemReveal was specifically designed to target this market
  -- Excellent for small spots, repeatable sample targeting

## Slide 17: Example 2: Quantitative Analysis on Aluminum and Steel
25 known samples
  23 Al alloy standards from ARMI and Alcoa
  2 measured at TSI using Arc/Spark-OES (9999, 9998)
Major elements Si, Cu, Mg, and Zn are studied

## Slide 18: Sampling and Data pre-processing
Sampling
  1064 nm, 160 mJ per shot, 200 μm spot size
  9 locations on a 3x3, 500 μm spaced grid
  25 cleaning shots, followed by 25 shots per spot
  4 channel spectrometer, 7545 variables, 180-900 nm
Data pre-processing
  SVG filter, background removal, Al 310 nm internal standard, spectral range 190-365 nm
  Leave-One-Sample-Out, and Root Mean Square Error of Cross Validation (RMSECV)
  Bias and Variance calculated (ANOVA)
  Matlab/ChemLytics for models and figures

## Slide 19: Spectral Range Selection on ChemLytics
Cu
Al
Mg
190-365 nm, 3700 variables

## Slide 20: Aluminum Univariate Calibration
Single or summed area of peaks
Linear at low concentrations, self-absorption at higher concentrations
Cu
Notes: nonliear regression, or "linear regression in nonlinear space..."
this is very standard LSR

## Slide 21: Aluminum Univariate Calibration
Good R2 and model fit
Some bias
Si
Mg

## Slide 22: Aluminum Multivariate
Cu
5 loading vectors
Supporting evidence: regression vector corresponds to element features.
Linear model with ability to partly model nonlinear features.
Univariate
Cu

## Slide 23: Aluminum samples: PLS
Good R2 and model fit
Improved statistics overall

## Slide 24: Example 2b: Typical LIBS Univariate calibration – Stainless Steel
Typically the best RSDs on homogenous materials ~1%
Data from Stipe, Buckley et al.,
Applied Spectroscopy 64 (2) p154. (Feb 2010)

## Slide 25: Example 2b: Stainless Steel Multivariate Analysis
Comparison of predicted versus actual concentrations
  Linear, slope = 1, intercept ≈ 0, R2 ≈ 1
Completed with customer samples
 | Slope | Intercept | r2 | RMSEP | Range (%)
Si | 1.00 ± 0.01 | 0.001 ± 0.005 | 0.9994 | 60 ppm | 0.01 - 0.62
P | 0.99 ± 0.03 | 0.0002 ± 0.0003 | 0.997 | 5 ppm | 0.001 - 0.023
S | 0.99 ± 0.01 | 0.003 ± 0.003 | 0.9993 | 43 ppm | 0.004 - 0.36
Cr | 1.00 ± 0.01 | -0.1 ± 0.3 | 0.999 | 487 ppm | 17.0 - 21.3
Mn | 1.00 ± 0.01 | 0.00 ± 0.01 | 0.9994 | 157 ppm | 0.01 - 1.62
Fe | 1.01 ± 0.01 | -0.5 ± 0.9 | 0.9991 | 442 ppm | 66 - 70
Ni | 1.00 ± 0.01 | 0.0 ± 0.1 | 0.999 | 453 ppm | 8.2 - 11.8
Cu | 0.99 ± 0.01 | 0.00 ± 0.02 | 0.9991 | 378 ppm | 0.1 - 3.2
Mo | 1.01 ± 0.02 | -0.002 ± 0.003 | 0.998 | 56 ppm | 0.01 - 0.32

## Slide 26: Example 3: LIBS versus XRF Results: Geological Samples
Mg
Fe
Al
Ca

## Slide 27: Depth profiling and surface mapping
This combination is a unique feature of LIBS
Depth profiling
  Semi-quantitative
  Depth per shot needs calibration for each material
Surface mapping can be made quantitative

## Slide 28: Depth profiling with LIBS
As laser pounds into material there is typically some influence from upper layers
Plasma goes into the crater
Side walls influence spectra slightly
Result: continued slight presence of elements from upper layers as depth increases
Laser
Laser
Laser
Laser
Successive laser shots

## Slide 29: Sample before ablation

## Slide 30: Sample after ablation
300 micron spots

## Slide 31: Depth Profiling: Sample 1

## Slide 32: Sample 2

## Slide 33: Sample 3

## Slide 34: LIBS Applications: Elemental Mapping
Elemental maps can be generated
Resolution determined by laser spot size
No sample cell required
2024 aluminum sample from www.onlinemetals.com
20x20 map produced with approximately 75 micron spots
Argon purge on sample

## Slide 35: Applications summary
LIBS capabilities were demonstrated for:
  Microanalysis
  Bulk analysis
  Complex samples
  Depth profiling / surface mapping
Quantification of elements comparable to other established elemental techniques
Calibration curves are material-specific

## Slide 36: Poll question #2
Which characteristic of LIBS would be most frequently used in your applications?
Quantification of major and minor elements in solid materials
Qualitative, rapid identification of materials and elements

## Slide 37: Chemreveal LIBS desktop Elemental analyzer

## Slide 38: ChemReveal Designed for Industrial & Research Applications
Industrial QA/QC Capable
  Technician-User Interface
  Higher-throughput analyses
  Designed for serviceability & Robustness
Research ready
  Detailed Analyst –level software
Fail
Pass
Pass
Pass

## Slide 39: ChemReveal Designed with Quality
  + TSI is an ISO9001:2008 QMS manufacturer
  + Products go through V&V and are supported by TSI’s global sales and service organization
  + Standards shipped with every
  ChemReveal Instrument

## Slide 40: Features of the ChemReveal LIBS Desktop Elemental Analyzer
Hardware:
Integrated system for conducting LIBS
Small footprint (24” x 20”)
Large sample chamber (9” x 5” x 5” samples)
XYZ stages with 2” travel and < 1 um resolution
CE Certified, Class I Laser product built under ISO9001:2008 quality management system

## Slide 41: Features of the ChemReveal LIBS Desktop Elemental Analyzer
Software:
Intuitive software for operation, method development and calibration
Software for different user levels
Report generation for record keeping
Notes: Peak ID, to peak and background definition, to calibration curve generation and quantification.

## Slide 42: Using the ChemReveal LIBS Desktop Elemental Analyzer
Sample preparation/loading:
Open door to sample chamber, and place sample inside
Close door and adjust sample to location of interest
Benefits:
No sample preparation
No vacuum
Fast loading and unloading of samples
Protected optics
Exhausted sample chamber
Notes: Purge gas used to protect optics and provide for higher sensitivity analysis of N, O, C, S.

## Slide 43: Running Methods using the Technician-mode Interface
Pass
12 | 3 | 5 | 2
1
2
3
4
Load check standard
Select Check Standard Method
Label Sample & Run
Get Pass/Fail and Quantification Results

## Slide 44: Workflow from Analyst to Technician
Raw LIBS
OES spectrum
Fail
Pass
Pass
Pass
Chart: Wavelength (nm)
35.5%
29.5%
35.2%
34.9%
Interpreted Data
Actionable Data
LIBS of solid samples
1. Robust LIBS
Consistent Data
Processing & Method
Development
3. Reproducible
Calibrations
4. Ease of Use

## Slide 45: ChemLytics Software for Method Development
Calibration & Method Development:
Peak area definition and integration
Univariate calibration model development (std)
Multivariate calibration model development (option)
All models created can be imported into a technician-mode interface
Notes: Amount of laser ablation is critical to the quantitation of any LIBS system. We control the laser power density to ensure this is consistent.

## Slide 46: ChemLytics Software for Method Development
Calibration & Method Development:
Peak area definition and integration
Univariate calibration model development (std)
Multivariate calibration model development (option)
All models created can be imported into a technician-mode interface
Notes: Amount of laser ablation is critical to the quantitation of any LIBS system. We control the laser power density to ensure this is consistent.

## Slide 47: ChemLytics Software for Reports
Report Generation:
Pass/Fail and check standards can be customized
Calibration and Quantification Results Reported
Notes: Spectral match feature against a database

## Slide 48: ChemLytics Software for Reports
Report Generation:
Pass/Fail and check standards can be customized
Calibration and Quantification Results Reported
Notes: Amount of laser ablation is critical to the quantitation of any LIBS system. We control the laser power density to ensure this is consistent.

## Slide 49: ChemReveal Validation Tests
  + Intraday reproducibility
  + Day-to-day reproducibility
  + Unit-to-unit reproducibility
  + Calibration Accuracy vs. Arc/Spark OES
  + Development of a daily check standard for instrument performance
~2 minutes to run each sample for these validation tests

## Slide 50: Results: Intraday Reproducibility
SRMs | Cr 520.45:Fe | Ni 361.94:Fe | Cu 324.75:Fe | Mn 403.08:Fe
SS 254SMO | 0.6% | 0.6% | 0.1% | 2.3%
SS 301 | 2.3% | 1.3% | 31.6% | 0.4%
SS 303 Se | 0.7% | 0.4% | 1.7% | 1.3%
SS 304 | 1.7% | 0.4% | 1.3% | 1.9%
SS 309 | 0.4% | 0.6% | 1.3% | 0.9%
SS 310 | 1.1% | 1.0% | 1.7% | 0.9%
%RSD of Element:Fe (438.35) ratio
3864 ChemReveal #1a

## Slide 51: Results: Intraday Reproducibility
SRMs | Cr 520.45:Fe | Ni 361.94:Fe | Cu 324.75:Fe | Mn 403.08:Fe
SS 254SMO | 1.1% | 1.2% | 0.5% | 1.7%
SS 301 | 2.5% | 2.5% | 15.5% | 1.5%
SS 303 Se | 1.8% | 2.6% | 1.4% | 2.7%
SS 304 | 1.1% | 2.3% | 1.2% | 1.6%
SS 309 | 1.2% | 2.1% | 0.9% | 1.6%
SS 310 | 1.9% | 1.6% | 1.7% | 1.8%
%RSD of Element:Fe (438.35) ratio
3864 ChemReveal #2

## Slide 52: Results: Multiple Day Reproducibility
SRMs | Cr 520.45:Fe | Ni 361.94:Fe | Cu 324.75:Fe | Mn 403.08:Fe
SS 254SMO | 1.8% | 4.1% | 3.8% | 3.2%
SS 301 | 2.1% | 2.2% | 17.0% | 2.9%
SS 303 Se | 2.0% | 1.7% | 2.9% | 1.3%
SS 304 | 1.1% | 1.8% | 3.3% | 2.1%
SS 309 | 1.1% | 0.6% | 2.2% | 2.4%
SS 310 | 0.8% | 2.9% | 3.8% | 3.8%
%RSD of Element:Fe (438.35) ratio
3864 ChemReveal #1b

## Slide 53: Results: Multiple Day Reproducibility
SRMs | Cr 520.45:Fe | Ni 361.94:Fe | Cu 324.75:Fe | Mn 403.08:Fe
SS 254SMO | 1.7% | 1.3% | 1.4% | 1.2%
SS 301 | 1.9% | 0.4% | 7.6% | 2.1%
SS 303 Se | 3.4% | 0.9% | 0.9% | 2.1%
SS 304 | 2.0% | 1.1% | 1.8% | 2.2%
SS 309 | 2.6% | 1.1% | 0.9% | 1.7%
SS 310 | 1.2% | 1.3% | 1.6% | 1.1%
%RSD of Element:Fe (438.35) ratio
3864 ChemReveal #2

## Slide 54: Results: Calibration
Ni Calibration | Slope (m) | Y-Intercept (b) | Correlation Coeff. (R^2)
Day 1 | 0.0193 | 0.2252 | 0.9985
Day 2 | 0.0187 | 0.2273 | 0.9986
Day 3 | 0.0183 | 0.2345 | 0.9987
Day 4 | 0.0186 | 0.2321 | 0.9968
Day 5 | 0.0183 | 0.2274 | 0.999
%RSD | 2.2% | 1.7% | 0.1%
Calibration reproducibility over multiple days for Ni:Fe (438.35) ratio
3864 ChemReveal #2
Chart: Ni Concentration (%)
| 17.7 | 7.12 | 9.17 | 8.3000000000000007 | 12.05 | 20.2
Day1 | 0.56508 | 0.36251 | 0.40163 | 0.38982 | 0.45035 | 0.61716
Day2 | 0.55628 | 0.36423 | 0.39891 | 0.382 | 0.44567 | 0.60764
Day3 | 0.5627 | 0.3641 | 0.40613 | 0.38489 | 0.45103 | 0.6011
Day4 | 0.57061 | 0.3641 | 0.39998 | 0.38707 | 0.45534 | 0.59984
Day5 | 0.55261 | 0.361 | 0.39651 | 0.37844 | 0.44238 | 0.59848






## Slide 55: Results: Apply Cal
 | LIBS Predicted Ni Concentration | Arc Spark | % Accuracy | RMSEP
SS 303 | 8.20 | 8.38 | 98% | 3.2%
SS 304 | 7.87 | 7.90 | 100% | 3.4%
SS 316 | 10.21 | 10.35 | 99% | 2.6%
SS 321 | 9.43 | 8.88 | 106% | 3.0%
SS 347 | 8.80 | 8.85 | 99% | 3.0%
3864 ChemReveal #1b

## Slide 56: Results: Apply Cal
 | LIBS Predicted Ni Concentration | Arc Spark | % Accuracy | RMSEP
SS 303 | 8.50 | 8.38 | 101% | 5.0%
SS 304 | 7.85 | 7.90 | 99% | 5.3%
SS 316 | 9.85 | 10.35 | 95% | 4.0%
SS 321 | 8.76 | 8.88 | 99% | 4.7%
SS 347 | 9.62 | 8.85 | 109% | 4.7%
3864 ChemReveal #2

## Slide 57: Validation Test Summary
ChemReveal LIBS Desktop Elemental Analyzer demonstrated:
  Good intraday reproducibility, %RSD <~3%
  Good day-to-day reproducibility, %RSD <~5%
  Exception is when measurements made on elements near LIBS DL of 10-100ppm
  Calibration is reproducible over multiple days (<3%)
  Compared to Arc/Spark-OES, RMSEP for Ni <5%
A check standard can be routinely used to check LIBS instrument performance

## Slide 58: ChemReveal LIBS vs. other methods

## Slide 59: LIBS complements other elemental analysis methods
  Pre-screen for ICP – analyze unknown matrices quickly to determine optimal digestion procedures for ICP, and whether ICP is necessary
  Rapid soils analysis - Analyze light elements and soil nutrients rapidly to complement XRF
  Rapid glass analysis - control glass production processes quickly to complement more laborious ICP that is used for QA
  Metals research – create new libraries for specialty alloys to quantify bulk and trace metals, and complement Arc/Spark-OES

## Slide 60: Test your samples at TSI Laboratories
Demonstration labs established in USA, China, and Germany for your sample analyses
State-of-the-art ICP-OES, Arc/Spark-OES and other analytical techniques to validate and calibrate LIBS for your samples
Tell us about your elemental analysis needs, send us your samples and get results

## Slide 61: Conclusions
LIBS is a rapid and versatile method for the analysis of solids
ChemReveal LIBS delivers a high-quality LIBS, designed for both research and industrial applications
Successful LIBS examples on glass, metals, and geochemical samples demonstrated elemental analysis, mapping and profiling
See what ChemReveal LIBS can do for your solid samples

## Slide 62: Acknowledgements
  Markus Gaelli
  Rob Robinsky
  Dahu Qi
  Dan Jensen
  Gregg Lithgow
  Kregg Philpott
  Alexia Fischer
Thank you for attending
Call D&N Scientific 614-270-9880
danw@columbus.rr.com