# Getting 12-bits of Resolution out of your 10-bit A/D

Contents

Simple Techniques that Lower Cost and Increase A/D Converter Resolution	1

Getting 12-bits of Resolution out of your 10-bit A/D - Part 1	1

Example Applications	1

Part 1: Increasing A/D Converter Resolution by Averaging	2

Averaging Technique	2

Temperature Example	3

Demonstration	3

Summary	4

Getting 12-bits of Resolution out of your 10-bit A/D - Appendix A	5

Part 1 Equations	5

Effects of Averaging on Quantization Noise	6

Practical Considerations	7

Getting 12-bits of Resolution out of your 10-bit A/D - Appendix B	9

Arrow Evaluation / Demonstration Tool	9

## Simple Techniques that Lower Cost and Increase A/D Converter Resolution

### Getting 12-bits of Resolution out of your 10-bit A/D - Part 1

Emerging products in smart sensors and instrumentation are demanding that engineers look for more ways to increase performance while simultaneously squeezing out cost. This article demonstrates how engineers can accomplish this by implementing digital signal processing techniques in common 8/16-bit microcontrollers. The reader will be able to increase the effective resolution of their own A/D designs after achieving a practical understanding of the algorithms used and seeing them demonstrated using an Arrow Electronics signal processing analysis tool and common microcontroller development platform.

### Example Applications

In many cases, designers select an A/D by initially considering the full scale range and minimum resolution. For example, a designer that is measuring temperature may have a resolution requirement of 0.1 degrees Fahrenheit over a 100 degree full scale range. Taking the full scale range and dividing it by the desired resolution implies a dynamic range of 1000. The dynamic range can also be expressed in decibels by taking 20 times the Log of 1000 which in this case is 60db. A pretty good approximation for the number of required A/D bits is to take the dynamic range in db, divided it by 6 and then round up to the nearest integer which implies that 10-bits are required for the A/D. However, with the implementation of some simple averaging algorithms in the microcontroller, the A/D requirement is easily reduced to 8-bits. Since many low-cost microcontrollers have 8-bit A/Ds in the peripheral set, the ability to fit your application into these solutions enables you to choose the lowest possible cost for your design.

The temperature measurement example focused on an application where the signal bandwidth (BW) was assumed to be very close to DC, however this is not always the case. For example, a design that monitors 60Hz sine-waves must maintain adequate bandwidth for that frequency in the data processing chain. If a minimum dynamic range of 70db is required, then an A/D with 12 bits would be chosen since 70 divided by 6 rounds up to 12. However, using the oversampling and moving average techniques presented in this article, an A/D with less than 12 bits could provide 70db of dynamic range or a 12-bit A/D already in the system could provide even better performance than 70db without sacrificing bandwidth.

This article addresses separately each of the previous examples for averaging and oversampling. The signal processing algorithms are presented in an overview manner along with the underlying assumptions that enable them to operate properly. Appendices are referenced for the reader who wants an explanation of how the presented equations, algorithms and demonstrations were developed.

### Part 1: Increasing A/D Converter Resolution by Averaging

Recall from our temperature measurement example that we needed 10 bits for the requirement of .1 degree resolution over a 100 degree full scale range. One solution would be to purchase a 10-bit A/D and be done with it! However, in situations where circuit cost is critical or when you're locked into a set number of bits because of legacy issues, considering some simple averaging techniques along with the hardware requirements achieves the best overall solution.

### Averaging Technique

So how does this actually work! Remember that an A/D's noise (QN) due to quantization error is related to the minimum A/D resolution according to Equation 1,

QN=Q23	Equation 1

where the minimum resolution, or quantization level (Q), is calculated by using Equation 2 and knowing the full scale range (FS) and the total number of bits (N) in the A/D converter.

Q=FS2N	Equation 2

Since quantization noise is directly proportional to the resolution of the A/D, processing A/D samples in a manner that reduces the noise, QN, can increase the A/D's effective resolution. Equation 3 describes what happens when you average a given number of samples (NS) from an A/D converter and then look at the quantization noise for that set of averaged samples (AQN). The effect is to reduce the quantization noise by the square-root of the number of samples averaged.

AQN=QNNS	Equation 3

For example, consider the algorithm defined by Equation 4 which is often called a moving average filter.

yn=1NSi=1NSxn−i	Equation 4

If we let x(n) represent A/D samples and y(n) represent processed samples, then the noise associated with each A/D sample would be QN and the noise associated with each averaged sample would be AQN. Using Equations 1 and 3, the averaged samples, y(n), will have lower noise and as a result lower quantization level (better resolution) than the original A/D samples.

An enabling assumption for this technique to work is that the A/D's RMS input noise must be random, zero mean, and greater than the quantization noise of the A/D. If this assumption isn't met, the averaged samples could "flat-line" causing an error to occur in the mean values. This can be overcome by using a summing amplifier which adds random noise to the input signal. Sometimes this can be a less expensive solution than increasing the number of bits in an external A/D. In many cases, the environment that the final system operates in guarantees the RMS noise to be greater than QN. If the mean value of the noise is not exactly zero, then a bias will be introduced into the data that can be easily calibrated out.

An equation can be derived that calculates the equivalent number of A/D bits (ENOB) for the averaged samples, y(n), as a function of the number of samples averaged, NS, the number of A/D bits, N, and the ratio of the A/D's RMS input noise to quantization noise, QN. If we call this ratio the Input Noise Factor (INF), then the moving average ENOB, y(n), can be calculated using Equation 5.

ENOB=N+1.66logNS−3.32logINF	Equation 5

For example, if the actual RMS input noise to the A/D is equal to QN, then INF is equal to 1. If it is 1.5 times QN, then INF would equal 1.5. Based on our previous assumptions, the value of INF must be greater than 1 for this technique to work. Practically, INF should be at least 1.5 or greater. Assume that we have an 8-bit A/D with RMS input noise equal to 1.5 times the quantization noise. If we let the number of samples averaged, NS, equal 64, then Equation 5 predicts an ENOB of just over 10 bits.

ENOB=8+1.66log64−3.32log1.5=10.41	Equation 6

Notice also that if we used a 10 bit A/D with INF still equal to 1.5, the new ENOB would be just over 12 bits. Please refer to Appendix A for a more detailed development of these equations and the assumptions used to derive them.

### Temperature Example

Getting back to our original temperature example: Let's assume that you scale your measurements into a 0 to 5V range and anticipate an RMS noise level on the order of 10mV. An 8-bit A/D will have a quantization level of 19.5mV and quantization noise of 6mV RMS. Using Equation 5, the INF is 10/6 or 1.67, and the ENOB would be

ENOB=8+1.66log64−3.32log1.67=10.26	Equation 7

or just over 10 bits assuming we averaged 64 samples. All things equal, it is best to average as many samples as possible if you have the available processing power and time to make the measurement. Assuming that your microcontroller has a 10-bit A/D, then Equation 5 would predict just under 12 bits of resolution for the same number of averaged samples.

### Demonstration

Figure 1 plots averaged samples taken from an 8-bit A/D where the analog DC input was shifted by 1/4 and 1/2 of the A/D's quantization level (Q/4, Q/2). The data was processed using Equation 4 and formatted into a 16-bit word. Both the average value and the RMS noise of the averaged samples are compared with what would be predicted by Equation 3. The RMS noise is plotted on the left axis and the mean values are plotted on the right axis. Appendix A describes the problem of not averaging enough samples and identifies 32 as a good minimum number. Figure 1 demonstrates this with the agreement between real-data and what is predicted by Equation 3 at 32 averages per sample. At 128 averages per sample, the Q/4 and Q/2 sub-quantization steps are very clearly defined. Appendix B identifies the actual development board hardware and software tool that was used to demonstrate this technique. The software that was created for the development board as well as the software evaluation tool can be obtained through your local Arrow FAE (or the Arrow Website).

Figure 1. Demonstrated Increase in Resolution for an 8-bit A/D

### Summary

Part 1 presented a practical understanding of how you can increase the resolution of an A/D using a moving average filter. We assumed that the noise at our A/D input was random, greater than the quantization noise of the A/D, and equally distributed about the mean of the parameter that we were measuring. Finally, we assumed that the measured parameter would not change significantly over the sampling interval. The technique was demonstrated using a hardware demonstration board and the Arrow signal processing analysis tool. Appendices A and B are available for the reader who wishes to have more detail on the development of the equations.

## Getting 12-bits of Resolution out of your 10-bit A/D - Appendix A

This appendix offers a more detailed development of the equations presented in Part 1 of the article: "Simple Techniques that Lower Cost and Increase A/D Converter Resolution".

### Part 1 Equations

The smallest increment that a single sample from an A/D converter can resolve is called the quantization level (Q). Equation (1) relates Q to the full scale range (FS) and number of bits (N) in the A/D converter.

Q=FS2N	Equation 8

Obviously, since the input to the A/D is continuous there will be some difference, QE, between the actual input and value assigned to the quantized output code. If we only consider this difference, we can calculate a value that closely approximates the root-mean-squared (RMS) error due to quantization. We call that RMS error quantization noise (QN).

At this point, we need to make some assumptions about the quantization error that are reasonable and help simplify the math to something that is practical to use. Figure 1 shows a single A/D code centered between the trip point of two adjacent codes. Since the actual input could be anywhere in the cross hatched area, the maximum QE that we would expect to see for any A/D code is Q/2 or "Q/2.

Figure 2. Quantization Error (QE)

Making this assumption ignores the A/D's differential non-linearity (DNL) and integral non-linearity (INL) specifications which indicate how consistent the value of Q is across the entire A/D range. In most cases, A/Ds have very good DNL and INL specifications so ignoring these variations is a reasonable assumption to make. In addition to the maximum values for QE, we also assume that all errors between Q/2 and "Q/2 are random and have equal probability of occurring. This means that the probability density function of QE between Q/2 and "Q/2 is uniform and equal to 1/Q. For further reading on probabilities and probability density functions (PDF), see the references at the end of this article.

With these assumptions in hand, we can easily derive an expression for A/D converter noise due to quantization error. The expected or mean value of the random variable QE can be written as:

QE=−Q/2Q/2QEPDFQEdQE	Equation 9

which is nothing more than the random variable QE times it's probability density function integrated over it's min/max excursions. Using our assumption that the PDF(QE) is uniform and equal to 1/Q, and then working through the integral gives a result that the mean value of the quantization error is zero (3).

QE=1Q−Q/2Q/2QEdQE=1QQE22−Q/2Q/2=0	Equation 10

Figure 2 demonstrates that the expected value of QE is in the center of the PDF(QE) over the interval "Q/2 to Q/2. For this interval, the PDF is centered about zero. Having zero mean is critical since it says that our assumptions about the A/D's QE do not bias the measurements that we are making with the A/D.

Figure 3. Probability Density Function (PDF) of QE

Having looked at the mean value of the random variable QE, we can also derive an expression for quantization noise, QN. We have assumed that QE is random with zero mean and uniformly distributed over the interval "Q/2 to Q/2. Equation (4) gives the general expression for the variance of QE where

VarQE=QN2=−Q/2Q/2QE−QE2PDFQEdQE	Equation 11

Letting the mean value of QE be equal to zero and replacing the PDF(QE) with 1/Q gives the resulting expression for mean squared error (5), and RMS noise, QN (6).

QN2=1Q−Q/2Q/2QE2dQE=1QQE33−Q/2Q/2=Q212	Equation 12

QN=Q23	Equation 13

Simply stated, the expected RMS noise level due to A/D quantization is a little less than 1/3 the quantization level.

### Effects of Averaging on Quantization Noise

An interesting question with respect to Equations 5 and 6 is what would happen to QN if we simply averaged the quantization error over time? For example, instead of looking at QE for each sample out of the A/D, we create a new random variable, AQE, that averages an arbitrary number of samples (NS) of the random variable QE as shown in Equation (7).

AQE=1NSn=1NSQEn	Equation 14

At this point, it is helpful to remember the original assumptions that we made about QE being uniformly distributed and random. If QE is random then individual values of QE are statistically independent meaning any correlation between different samples would be zero. This allows us to calculate the variance of AQE by summing the individual variances and ignoring the correlation terms that result when you add multiple noise sources together. Equations 8 and 9 take the variance of AQE in Equation 7 by using two properties:

1. The variance of a sum of independent variables is equal to the sum of variances, Var(X+Y) = Var(X) + Var(Y), and 2. The variance of a constant times a random variable is equal to the square of that constant times the variance, Var(cX) = c2Var(X).

Therefore,

VarAQE=AQN2=n=1NSVar1NSQEn=1NS2n=1NSVarQEn=QN2NS	Equation 15

AQN=QNNS	Equation 16

which simply says that when we average samples, the quantization noise will be reduced by the square root of the number of samples. This can be used to calculate a new effective number of bits (ENOB) for the original A/D. For example, NS = 16 would reduce the noise by a factor of 4. According to Equation (6) this reduces the effective quantization level by a factor of 4 and according to (1) increases the effective number of bits by a factor of 2. Using the same analysis, Equation 10 is derived and used to approximate the ENOB as a function of the number of samples averaged and the ratio of actual RMS input noise to QN which we will call the Input Noise Factor (INF). For example, if the RMS input noise is QN, then INF is equal to 1. If it is 1.5 times QN, then INF would equal 1.5. Remember that INF must be greater than 1 for this technique to work and practically should be at least 1.5 or greater. So, for example, if we have an 8-bit A/D with RMS input noise equal to 1.5 times the quantization noise, then averaging 128 samples would produce samples with an ENOB of just under 11 bits.

ENOB=N+1.66logNS−3.32logINF	Equation 17

### Practical Considerations

Obviously there are some practical considerations when using this technique. First, we have assumed the presence of random system noise that is greater than the RMS quantization noise. If this assumption isn't met, the averaged samples could "flat-line" causing a bias to occur in the mean value. This can be overcome by using a summing amplifier at the A/D input which adds in random noise generated by some external source or pulse-width-modulated (PWM) output from the local microprocessor. Sometimes this can be a much cheaper solution than increasing the number of bits in an external A/D converter. Obviously, if the A/D is located on-chip then changing it's physical resolution isn't an option so this may be a way to increase the accuracy of your measurements when backed into that corner. In many cases, the environment that the final system will operate in guarantees the RMS noise to be greater than QN. In that situation, this technique can be used without a summing amplifier being required, assuming that the noise is reasonably random.

Second, recall that the mean value and noise were derived using an assumed uniform, random PDF(QE). If the mean shifts slightly, so will the PDF(QE). Figure 4 illustrates a practical way to visualize this. Consider an A/D input centered between two codes. If our assumptions about the PDF(QE) are correct, then we will see an even number of A/D codes above and below the actual input signal. The mean of those codes would correctly predict the actual input to be centered between the two codes. If the mean shifts by Q/4, then we would expect to see a corresponding shift in the PDF(QE) and proportional change in the number of codes above and below the actual input. The calculation of the mean would again correctly predict the actual shifted input because of the integration of the PDF(QE) over the A/D intervals and redistribution of samples that fall into each of the A/D code intervals. For example, recall the definition of mean value in Equation (2). If the two A/D codes in Figure 4 happened to be 1024 and 1025, then the mean would be 1024(.5) + 1025(.5) or 1024.5. For the shifted input, 1024(.25) + 1025(.75) would give a mean of 1024.75. It can be shown that the same result is produced by simply averaging the A/D samples. This assumes that we have averaged enough samples to accurately reflect the distribution of the PDF into the different A/D code intervals. Typically, a minimum of 32 samples is required to get a statistically significant sampling under the assumptions that we have made, however look through the references if you want to drill down further into how this number is arrived at.

Figure 4. Discrete PDF as Identified by A/D Codes

In reality, the input noise will likely have a normal Gaussian distribution which only serves to help our situation further since the average of normally distributed samples will produce outputs that are also normally distributed. Also, the RMS of a normal distribution is its standard deviation (sigma). We can assume that ~99% of the noise spikes will be within 3-sigma and evenly distributed about the mean of the input signal. The PDF(QE) will track shifts in the mean and can be identified by the proportional redistribution of A/D codes about the true mean value.

## Getting 12-bits of Resolution out of your 10-bit A/D - Appendix B

This appendix offers a practical method for demonstrating the concepts presented in the article: "Simple Techniques that Lower Cost and Increase A/D Converter Resolution". The demonstrations use software running on either of two Motorola demo boards (908QT4 or 9S08GB60) and an Arrow signal processing analysis tool.

Experimenting with Real Data was collected using the latest Motorola demo boards for the 908QT4 and new 9S08GB60 processors. These demo boards include microprocessors with 8/10 bit A/Ds on the chip. Figure 1 shows a block diagram of the configuration that was used to take the data. The Arrow analysis tool interfaces with code on the microprocessor demo board target. The target code samples data from the A/D, averages it, then formats it into a 16-bit word and transfers it out to the analysis tool via the serial link. The RMS noise at the A/D input was on the order of ½Q which is right at the lower limits for this technique. Because the data is formatted into a 16-bit word, 256 counts would correspond to Q and QN would be equal to 74 counts.

Figure 5. Motorola Evaluation Tool Configuration

### Arrow Evaluation / Demonstration Tool

Figure 2 is a screenshot of the tool used to collect samples from the target and demo the concepts discussed in the article. The Arrow tool was implemented using Excel and Visual Basic for Applications. It interfaces with the Motorola target board using an RS-232 asynchronous serial interface running at 9600 baud. The Excel program transmits the number of samples to average to the target demo board and then collects the data in a pre-defined packet that is displayed by the tool. The number of samples processed is defined by the tool's general purpose user interface (GUI). The target software collects A/D samples, averages them, formats them into a 16-bit word and then transfers them back via the serial interface to the GUI program. Statistics are collected and then time plots and histograms of the sampled data points are displayed. For a more detailed description of the complete Arrow signal processing tool and to download the code that runs on the Motorola demo boards, contact your local Arrow FAE.

Figure 6. Arrow Electronics Evaluation / Demo Tool for Increasing A/D Resolution